Direct Ophthalmoscope

The direct ophthalmoscope is commonly used for routine examination of the fundus of the eye, especially when a slit lamp cannot be used. It is small, easily portable and can also be used to examine the more anterior parts of the eye.

The instrument consists of a system of lenses which focus light from an electric bulb on to a mirror where a real image of the bulb filament is formed.

The mirror reflects the emitted light in a diverging beam which is used to illuminate the patient's eye. The mirror contains a hole through which the observer views the illuminated eye.

The area of retina which can be seen at any one time is called the field of view. It is governed by the projected image of the sight-hole on the retina.

the field of view is considerably enlarged when the pupil is dilated; hence the advantage of instilling a mydriatic prior to fundoscopy.

In order to utilise the maximum available field of view it is necessary for the observer to be as close as possible to the patient's eye.

the image size varies with the refractive state of the patient's eye, the image being smaller in hypermetropia, and larger in myopia than in emmetropia.

In order to see a clear view of the patient's retina in myopia or hypermetropia, it is therefore necessary to use a correcting lens to bring the image to a focus on the observer's retina. A range of correcting lenses is incorporated in the ophthalmoscope.

For those observers who have a refractive error there are two possibilities. The observer can remove his spectacles and rack up the appropriate lens in the ophthalmoscope to give a clear view of the patient's fundus. The appropriate lens is the algebraic sum of his own and the patient's refractive error. Alternatively, he can use the instrument with his glasses on. However, his field of view will be restricted as the sight-hole in the mirror will be further from his eye.

Indirect Ophthalmoscope

The indirect ophthalmoscope offers an alternative means of examining the retina and vitreous, and gives a very different view of the retina compared with that obtained with the direct ophthalmoscope. Both instruments have their place and should be used to complement each other during the clinical examination of the eye.

In the indirect method of ophthalmoscopy a powerful convex lens (called the condensing lens) is held in front of the patient's eye. The usual powers used are +20 D and +13 D. The illuminating light beam passes through the condensing lens into the eye and light reflected from the retina is refracted by the condensing lens to form a real image between the condensing lens and the observer. The observer studies this real image of the patient's retina.

Illumination is usually provided by an electric lamp mounted on the observer's head. Light from this source is rendered convergent by the condensing lens. Thus a convergent beam enters the patient's eye and is brought to a focus within the vitreous by the eye's refractive system.

Summary of the optical properties of the direct and indirect ophthalmoscope.

|  | Direct ophthalmoscope | Indirect ophthalmoscope |
| --- | --- | --- |
| Image | Not inverted | Vertically and horizontally inverted |
| Field of view | Small (6°) | Large (25°) |
| Magnification | Large ( × 15) | Small ( × 3 [+20D]) ( × 5 [+13D]) |
| Binocularity | Not available | Stereoscopic view |
| Influence of patient's refractive error | Large | Small |
| Teaching facility | None | Teaching mirror |

The Relative Merits of the Direct and Indirect Ophthalmoscope

Other considerations include:

(1) Size of the instrument. The direct ophthalmoscope is much smaller and lighter than the indirect, and many models are pocket size.

(2) Illumination is crucial to the view obtained. Because indirect ophthalmoscopes are larger instruments, there is more scope for fitting more powerful light sources. This renders the indirect ophthalmoscope more useful in examining patients with opacities in the ocular media.

(3) The combination of good illumination and wide field of view makes the indirect ophthalmoscope theinstrument of choice when examining patients with retinal detachments. If there is extensive subretinal fluid and the illumination is poor, an underlying malignancy may not be seen with the direct ophthalmoscope. Furthermore, examination with the indirect ophthalmoscope also allows identation of the peripheral retina.

(4) The indirect ophthalmoscope is also the ophthalmoscope of choice for use during retinal detachment surgery because it is used at a distance which allows the surgeon to preserve a sterile operative field.

(5) Laser energy can be delivered through the indirect ophthalmoscope to effect retinal photocoagulation.

Retinoscope

An accurate objective measurement of the refractive state of an eye can be made using the retinoscope.

The technique is called retinoscopy.

Light is directed into the patient's eye to illuminate the retina (the illumination stage). An image of the illuminated retina is formed at the patient's far-point (the reflex stage). The image at the far-point is located by moving the illumination across the fundus and noting the behaviour of the luminous reflex seen by the observer in the patient's pupil (the projection stage).

In hypermetropia the luminous reflex seen in the patient's pupil moves in the same direction as the illuminating light – a 'with' movement is observed.

When the patient has myopia the luminous reflex appears to move in the opposite direction to the illuminating light; 'against' movement.

The point of reversal or neutral point of retinoscopy is reached when the patient's far-point coincides with the observer's nodal point.

The observer sees a diffuse bright red reflex. This is because the movement of the reflex is infinitely rapid.

In practice, lenses are placed in front of the patient's eye until the point of reversal is seen by the observer. A correction is made for the working distance (add –1.5 D for 2/3 m, add –1.0 D for 1 m) and the corrected value of the lenses equals the patient's refractive error.

Slit Lamp

The routine method of examining the outer segment of the eye is with a slit lamp. It consists essentially of a relatively low-powered binocular compound microscope  which is linked to an adjustable bright light source. The instrument is known as a slit lamp because in everyday use the illumination is arranged so that a narrow vertical slit of light is projected on to the eye.

There are several different types of slit lamp but they all share certain basic features. The alignment of the microscope and the illumination is such that the point on which the microscope is focused corresponds to the point on which the light is focused. This is achieved by the microscope and the lighting system having a common focal plane. Their common axis of rotation also lies in this focal plane.

Another common feature is that there is a considerable distance between the microscope itself and the patient's eye. This is because the microscopes have a long working distance. The resulting gap between the microscope and the patient allows the observer to carry out certain manoeuvres such as removing a foreign body from the cornea. It also gives room to interpose certain optical devices, for example, a +90 D lens or a three-mirror contact lens, which permit inspection of the vitreous and retina.

The microscope tubes themselves are shortened by the incorporation of prisms, which also invert the image vertically and horizontally so that it appears erect and the right way round to the observer. A bank of Galilean telescopes of different powers is incorporated in most slit lamps to allow the magnification of the observation system to be varied.

Filters

The illumination system of many slit lamps incorporates various filters. The blue cobalt filter is used during applanation tonometry and so is commonly employed. Both the blue and the green (red-free) filters are useful in examination of the vitreous.

Fundus Examination

With the basic slit lamp it is not possible to see further back into the eye than the anterior third of the vitreous. This is because the refractive power of the cornea and lens renders light emerging from the deeper points of the eye parallel. Therefore no image is formed within the focal range of the slit lamp microscope. However, the fundus can be rendered visible by using an additional lens to overcome the refractive power of the eye.

90 D and 78 D Lenses

The principle of the indirect ophthalmoscope has been adapted so that the real image of the retina formed by the condensing lens  may be viewed through a slit lamp microscope.

For this purpose high power condensing lenses are used, 90 D or 78 D in order to shorten the light path and bring the retinal image within the focal range of the slit lamp.

The loss of image size through using such high power condensing lenses is more than compensated for by the magnification of the slit lamp microscope. The 90 D lens gives a wider field of view but less magnification than the 78 D lens.

The technique gives an excellent view of the posterior pole of the fundus but is less suitable for examining the peripheral retina.

Applanation Tonometer

The applanation tonometer is used to measure the intraocular pressure. The tonometer head is applied to the cornea with sufficient force to produce a standard area of contact. The force required is directly proportional to the intraocular pressure when the area of contact is 3.06  mm in diameter.

In order to achieve a standard area of corneal contact the applanation head contains two prisms, mounted with their bases in opposite directions.  The operator thus looks through the applanation head and sees the circle of corneal contact split into two half circles which are laterally displaced in opposite directions by the prisms. The operator adjusts the applanation pressure until the half circles just overlap one another when the area of contact will be exactly 3.06 mm in diameter. (The inner edges of the fluid meniscus define the area of contact and are aligned as shown.)

Non-Contact Tonometer

Tonometers have been developed which flatten the cornea with a puff of air, avoiding the need for the instrument to touch the eye.

Corneal applanation is measured by collecting light reflected from the central cornea. A parallel beam of light is directed on to the central cornea at an angle of 30° and the reflected light is measured by a photodetector at an angle of reflection of 30°.

The reflected beam of light will be strongest at this angle when the cornea is flat and acting as a plane mirror, rather than as a curved mirror. The instrument records the force of air required to flatten the cornea and displays the intraocular pressure which corresponds to that force.